## Summary
¦ | Scenario | ¦ | Description | ¦ | Good | ¦ | Decent | ¦ | Bad | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | AsthmaAttackModerateAcute | ¦ | Moderate acute asthma attack | ¦<span class="success"> | 18 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | AsthmaAttackSevereAcute | ¦ | Severe acute asthma attack | ¦<span class="success"> | 13 | </span>¦<span class="warning"> | 2 | </span>¦<span class="danger"> | 3 | </span>¦
¦ | AsthmaAttackLifeThreateningAcute | ¦ | Life threatening acute asthma attack | ¦<span class="success"> | 17 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 1 | </span>¦
¦ | COPDSevereEmphysema | ¦ | GOLD Stage III Emphysema | ¦<span class="success"> | 3 | </span>¦<span class="warning"> | 6 | </span>¦<span class="danger"> | 1 | </span>¦
¦ | COPDSevereBronchitisLeft | ¦ | Severe Chronic Bronchitis | ¦<span class="success"> | 4 | </span>¦<span class="warning"> | 5 | </span>¦<span class="danger"> | 1 | </span>¦
¦ | LobarPneumoniaSevereLeftLobe | ¦ | Severe Lobar Pneumonia in one lobe in the left lung | ¦<span class="success"> | 5 | </span>¦<span class="warning"> | 2 | </span>¦<span class="danger"> | 2 | </span>¦
¦ | LobarPneumoniaSevereRightLung | ¦ | Severe Lobar Pneumonia in two lobes of right lung | ¦<span class="success"> | 7 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 2 | </span>¦
¦ | LobarPneumoniaModerateBothLungs | ¦ | Moderate Lobar Pneumonia in both lungs | ¦<span class="success"> | 6 | </span>¦<span class="warning"> | 1 | </span>¦<span class="danger"> | 2 | </span>¦
¦ | TensionPneumothoraxOpenVaried | ¦ | Varied open pneumothorax severities and interventions | ¦<span class="success"> | 32 | </span>¦<span class="warning"> | 4 | </span>¦<span class="danger"> | 6 | </span>¦
¦ | TensionPneumothoraxClosedVaried | ¦ | Varied closed pneumothorax severities and interventions | ¦<span class="success"> | 25 | </span>¦<span class="warning"> | 4 | </span>¦<span class="danger"> | 6 | </span>¦
¦ | TensionPneumothoraxBilateral | ¦ | Open and closed bilateral pneumothorax and interventions | ¦<span class="success"> | 19 | </span>¦<span class="warning"> | 6 | </span>¦<span class="danger"> | 3 | </span>¦
¦ | AirwayObstructionVaried | ¦ | Airway Obstruction with varying severities | ¦<span class="success"> | 27 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 3 | </span>¦
¦ | Bronchoconstriction | ¦ | Bronchoconstriction with varying severities | ¦<span class="success"> | 24 | </span>¦<span class="warning"> | 4 | </span>¦<span class="danger"> | 2 | </span>¦
¦ | MainstemIntubation | ¦ | Right and left mainstem intubation and correction (with Succs) | ¦<span class="success"> | 22 | </span>¦<span class="warning"> | 3 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | EsophagealIntubation | ¦ | Esophageal intubation and correction (with Succs) | ¦<span class="success"> | 12 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 3 | </span>¦
¦ | Apnea | ¦ | Varied severities of respiratory apnea | ¦<span class="success"> | 6 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ |  | ¦ | Total | ¦<span class="success"> | 240 | </span>¦<span class="warning"> | 37 | </span>¦<span class="danger"> | 35 | </span>¦

## AsthmaAttack
AsthmaAttackModerate
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Trachea Flow - Peak Expiratory Flow (L/min) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Moderate acute asthma attack (Severity = 0.3) | ¦ |  | ¦ | 30 | ¦ | 550 | ¦<span class="success"> | 75% to 90% of normal @cite britishGuideline2014asthma | </span>¦<span class="success"> | Slight Increase @cite britishGuideline2014asthma | </span>¦<span class="success"> | ~95% of normal @cite limmer2011emergency | </span>¦<span class="success"> | No Change @cite britishGuideline2014asthma | </span>¦<span class="success"> | No Change @cite papiris2001clinical | </span>¦<span class="success"> | No Change @cite britishGuideline2014asthma | </span>¦<span class="success"> | No Change @cite britishGuideline2014asthma | </span>¦<span class="success"> | No Change @cite britishGuideline2014asthma | </span>¦<span class="success"> | Decrease @cite van1991physical | </span>¦
¦ | No asthma (Severity = 0.0) | ¦ |  | ¦ | 580 | ¦ | 775 | ¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦
AsthmaAttackSevere
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Trachea Flow - Peak Expiratory Flow (L/min) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Severe acute asthma attack (Severity = 0.7) | ¦ |  | ¦ | 30 | ¦ | 550 | ¦<span class="success"> | < 50% of Normal @cite papiris2001clinical, @cite britishGuideline2014asthma | </span>¦<span class="danger"> | > 25 @cite britishGuideline2014asthma | </span>¦<span class="success"> | < 60% of normal @cite limmer2011emergency | </span>¦<span class="danger"> | > 120 @cite papiris2001clinical; > 110 @cite britishGuideline2014asthma | </span>¦<span class="danger"> | Decrease Pulsus Paradoxus 85 mm Hg @cite papiris2001clinical | </span>¦<span class="warning"> | < 90% @cite papiris2001clinical; < 92% @cite britishGuideline2014asthma | </span>¦<span class="warning"> | Decrease Mild Hypoxemia @cite bergeronSME; < 60 mmHg @cite papiris2001clinical, @cite britishGuideline2014asthma | </span>¦<span class="success"> | Normal or Slight Increase > 45 mm Hg @cite papiris2001clinical, @cite britishGuideline2014asthma | </span>¦<span class="success"> | Decrease @cite van1991physical | </span>¦
¦ | No asthma (Severity = 0.0) | ¦ |  | ¦ | 580 | ¦ | 775 | ¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦
AsthmaAttackLifeThreatening
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Trachea Flow - Peak Expiratory Flow (L/min) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Life threatening acute asthma attack (Severity = 0.9) | ¦ |  | ¦ | 30 | ¦ | 550 | ¦<span class="success"> | < 33% of Normal @cite britishGuideline2014asthma | </span>¦<span class="success"> | > 25 @cite britishGuideline2014asthma | </span>¦<span class="success"> | < 60% of normal @cite limmer2011emergency | </span>¦<span class="success"> | > 120 @cite papiris2001clinical; > 110 @cite britishGuideline2014asthma Arrhythmia, @cite britishGuideline2014asthma | </span>¦<span class="danger"> | Decrease NO Pulsus Paradoxus 85 mm Hg @cite papiris2001clinical | </span>¦<span class="success"> | < 92% @cite britishGuideline2014asthma | </span>¦<span class="success"> | < 60 mmHg @cite britishGuideline2014asthma | </span>¦<span class="success"> | Normal or Slight Increase > 45 mm Hg @cite papiris2001clinical, @cite britishGuideline2014asthma | </span>¦<span class="success"> | Decrease @cite van1991physical | </span>¦
¦ | No asthma (Severity = 0.0) | ¦ |  | ¦ | 580 | ¦ | 775 | ¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦<span class="success"> | Normal | </span>¦

## COPD
COPDSevereEmphysema
¦ | Segment | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Trachea Flow - Peak Expiratory Flow (L/min) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦ | Aorta Lactate (mg/dL) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Severe Emphysema: Bronchitis Severity = 0.5 Emphysema Severity = 0.7 Both Lungs 100% | ¦ | GOLD Stage III | ¦ | 120 | ¦<span class="success"> | < 70% of Normal @cite perez2009can | </span>¦<span class="warning"> | Increase, Tachypnea, Dypsnea, > 30 @cite gunning2003pathophysiology | </span>¦<span class="warning"> | < 60% of normal @cite gunning2003pathophysiology | </span>¦<span class="success"> | Increase, Tachycardia | </span>¦<span class="danger"> | Increase, Pulmonary Hypertension , > 140 mm Hg @cite keller2003pathophysiology, @cite scharf2002hemodynamic | </span>¦<span class="warning"> | < 90% @cite man2003contemporary; < 89% @cite keller2003pathophysiology | </span>¦<span class="warning"> | Decrease, Hypoxemia @cite keller2003pathophysiology; < 55 mm Hg @cite man2003contemporary, @cite keller2003pathophysiology | </span>¦<span class="warning"> | Increase, Hypercapnia, > 55 mmHg @cite keller2003pathophysiology | </span>¦<span class="success"> | Decrease @cite van1991physical | </span>¦<span class="warning"> | Increase @cite mathur1999cerebral | </span>¦
COPDSevereBronchitis
¦ | Segment | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Trachea Flow - Peak Expiratory Flow (L/min) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦ | Aorta Lactate (mg/dL) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Severe Chronic Bronchitis: Bronchitis Severity = 0.65 Emphysema Severity = 0.5 Both Lungs 100% | ¦ |  | ¦ | 120 | ¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="success"> | > 20 @cite bergeronSME | </span>¦<span class="warning"> | < 60% of normal @cite bergeronSME | </span>¦<span class="success"> | Increase, Tachycardia @cite bergeronSME | </span>¦<span class="danger"> | Increase, Pulmonary Hypertension, > 140 mm Hg @cite bergeronSME | </span>¦<span class="warning"> | < 90% @cite bergeronSME | </span>¦<span class="warning"> | Decrease, Hypoxemia, < 55 mm Hg @cite bergeronSME | </span>¦<span class="warning"> | Increase, Hypercapnia > 55 mmHg @cite bergeronSME | </span>¦<span class="success"> | Decrease @cite van1991physical, @cite bergeronSME | </span>¦<span class="warning"> | Increase @cite mathur1999cerebral | </span>¦

## LobarPneumonia
LobarPneumoniaSevereLeftLobe
¦ | Condition | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦ | Core Body Temperature (Degrees C) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Severe Lobar Pneumonia: Severity = 0.70 Left Lung = 50% Right Lung = 0% | ¦ | Severe Lobar Pneumonia in one lobe in the left lung. Current model does not include metabolic effects. | ¦ | 120 | ¦<span class="success"> | Increase, Tachypnea, Dypsnea > 20 @cite ebell2007predicting | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="warning"> | Increase, Tachycardia, > 100 @cite ebell2007predicting | </span>¦<span class="danger"> | Decrease, < 90 mm Hg @cite fine1997prediction | </span>¦<span class="warning"> | Decrease , < 95% @cite majumdar2010oxygen | </span>¦<span class="danger"> | Decrease, Hypoxemia @cite fine1997prediction < 60 mm Hg | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="danger"> | Increase, > 37.8 C (100 F) @cite ebell2007predicting | </span>¦
LobarPneumoniaSevereRightLung
¦ | Condition | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦ | Core Body Temperature (Degrees C) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Severe Lobar Pneumonia: Right Lung Severity = 0.70 Left Lung = 0% Right Lung = 67% | ¦ | Severe lobar pneumonia in two lobes of the right lung. Current model does not include metabolic effects. | ¦ | 120 | ¦<span class="success"> | Increase, Tachypnea, Dypsnea > 20 @cite ebell2007predicting | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="warning"> | Increase, Tachycardia, > 100 @cite ebell2007predicting | </span>¦<span class="danger"> | Decrease, < 90 mm Hg @cite fine1997prediction | </span>¦<span class="success"> | Decrease , < 95% @cite majumdar2010oxygen | </span>¦<span class="warning"> | Decrease, Hypoxemia @cite fine1997prediction < 60 mm Hg | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="danger"> | Increase, > 37.8 C (100 F) @cite ebell2007predicting | </span>¦
LobarPneumoniaModerateBothLungs
¦ | Condition | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦ | PaO2 (mmHg) | ¦ | PaCO2 (mmHg) | ¦ | IERatio | ¦ | Core Body Temperature (Degrees C) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Moderate Lobar Pneumonia: Severity = 0.2 Left Lung = 100% Right Lung = 100% | ¦ | Moderate pneumonia in both lungs. Current model does not include metabolic effects. | ¦ | 120 | ¦<span class="success"> | Increase, Tachypnea, Dypsnea > 20 @cite ebell2007predicting | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="warning"> | Increase, Tachycardia, > 100 @cite ebell2007predicting | </span>¦<span class="danger"> | Decrease, < 90 mm Hg @cite fine1997prediction | </span>¦<span class="warning"> | Decrease , < 95% @cite majumdar2010oxygen | </span>¦<span class="danger"> | Decrease, Hypoxemia @cite fine1997prediction < 60 mm Hg | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="success"> | Decrease @cite bergeronSME | </span>¦<span class="danger"> | Increase, > 37.8 C (100 F) @cite ebell2007predicting | </span>¦

## TensionPneumothorax
TensionPneumothoraxOpenVaried
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Mean Arterial Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Tension Pneumothorax; severity = 0.0 | ¦ | A severity of zero should not change the leak resistance from the default open switch value. | ¦ | 0 | ¦ | 15 | ¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦
¦ | Tension Pneumothorax; severity = 0.3 | ¦ | Represents a "sucking" chest wound on the left lung side with a mild severity (medium size hole). | ¦ | 30 | ¦ | 300 | ¦<span class="success"> | No change or Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Tension Pneumothorax; severity = 0.6 | ¦ | Represents a "sucking" chest wound on the left lung side with a moderate severity (large size hole). | ¦ | 360 | ¦ | 600 | ¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Tension Pneumothorax; severity = 1.0 | ¦ | Represents a "sucking" chest wound on the left lung side with the severe severity (maximum size hole). Should lead to full collapse. | ¦ | 660 | ¦ | 900 | ¦<span class="success"> | Clinical sign: Tachypnea @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite leigh2005tension | </span>¦<span class="success"> | Clinical sign: Tachycardia @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Clinical sign: Hypoxia @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Apply Chest Occlusive dressing on the left side | ¦ | Represents the closing of chest wound on the left side | ¦ | 960 | ¦ | 1000 | ¦<span class="success"> | No Change or Increased @cite bergeronSME | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | Tachycardia > 120 @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Modest decrease to no change @cite bergeronSME | </span>¦
¦ | Apply needle decompression | ¦ | Needle decompression is applied on the left side of the chest | ¦ | 1020 | ¦ | 1320 | ¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="warning"> | Returning toward normal @cite bergeronSME | </span>¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="success"> | Normal @cite bergeronSME | </span>¦
TensionPneumothoraxClosedVaried
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Mean Arterial Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Tension Pneumothorax; severity = 0.0 | ¦ | A severity of zero should not change the leak resistance from the default open switch value. | ¦ | 0 | ¦ | 15 | ¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | No Change | </span>¦
¦ | Tension Pneumothorax; severity = 0.3 | ¦ | Represents a leak on the right lung side with a mild severity (medium size hole). | ¦ | 30 | ¦ | 300 | ¦<span class="success"> | No change or Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | No change or Decrease @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Tension Pneumothorax; severity = 0.6 | ¦ | Represents a leak on the right lung side with a moderate severity (large size hole). | ¦ | 360 | ¦ | 600 | ¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Tension Pneumothorax; severity = 1.0 | ¦ | Represents a leak on the right lung side with the severe severity (maximum size hole). Should lead to full collapse. | ¦ | 660 | ¦ | 900 | ¦<span class="success"> | Clinical sign: Tachypnea @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite leigh2005tension | </span>¦<span class="success"> | Clinical sign: Tachycardia @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Clinical sign: Hypotension @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Clinical sign: Hypoxia @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Apply needle decompression | ¦ | Needle decompression is applied on the right side of the chest | ¦ | 960 | ¦ | 1260 | ¦<span class="success"> | Slightly lower tachypnea @cite bergeronSME | </span>¦<span class="warning"> | Returning toward normal @cite bergeronSME | </span>¦<span class="success"> | Tachycardia, but may be lower value @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="success"> | Normal @cite bergeronSME | </span>¦
TensionPneumothoraxBilateral
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Tidal Volume (mL) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Mean Arterial Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Tension Pneumothorax; Severity: 0.5 Type: Open Side: Left Tension Pneumothorax; Severity: 0.5 Type: Closed Side: Right | ¦ | Represents a leak on the right and left lung sides simultaneously with a moderate severity (large size hole). | ¦ | 30 | ¦ | 300 | ¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Increase @cite leigh2005tension | </span>¦<span class="success"> | Increase @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="danger"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦<span class="success"> | Decrease @cite bond200968w, @cite leigh2005tension | </span>¦
¦ | Apply Chest Occlusive dressing on the left side | ¦ | Represents the closing of chest wound on the left side | ¦ | 330 | ¦ | 360 | ¦<span class="success"> | No Change or Increased @cite bergeronSME | </span>¦<span class="success"> | No Change | </span>¦<span class="success"> | Tachycardia > 120 @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Slightly improved to no change @cite bergeronSME | </span>¦<span class="success"> | Modest decrease to no change @cite bergeronSME | </span>¦
¦ | Apply needle decompression on the left side | ¦ | Needle decompression is applied on the left side of the chest | ¦ | 390 | ¦ | 400 | ¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="warning"> | Diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Diminished from normal @cite bergeronSME | </span>¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦
¦ | Apply needle decompression on the right side | ¦ | Needle decompression is applied on the right side of the chest | ¦ | 420 | ¦ | 720 | ¦<span class="success"> | Slightly lower tachypnea @cite bergeronSME | </span>¦<span class="success"> | Returning toward normal @cite bergeronSME | </span>¦<span class="success"> | Tachycardia, but may be lower value @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="warning"> | Slightly diminished from normal @cite bergeronSME | </span>¦<span class="success"> | Normal @cite bergeronSME | </span>¦

## AirwayObstructionVaried
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Trigger airway obstruction; severity = 0.3 | ¦ | Represents the partial (severity = 0.3) obstruction of airway due to foreign body (obstructing material). Not severe enough to trigger CV responses. | ¦ | 30 | ¦ | 180 | ¦<span class="success"> | Increase @cite Morgan2006Clinical | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="danger"> | Increase @cite bergeronSME(case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Increase @cite bergeronSME (case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Decrease @cite Morgan2006Clinical | </span>¦
¦ | Remove airway obstruction | ¦ | Airway obstruction is turned off. | ¦ | 210 | ¦ | 300 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦
¦ | Turn on airway obstruction; severity = 0.6 | ¦ | Represents the partial obstruction (severity = 0.6) of airway due to foreign body. | ¦ | 510 | ¦ | 180 | ¦<span class="success"> | Increase > Severity 0.3 @cite Morgan2006Clinical | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="danger"> | Increase @cite bergeronSME (case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Increase @cite bergeronSME (case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Decrease < Severity 0.3 @cite Morgan2006Clinical | </span>¦
¦ | Remove airway obstruction | ¦ | Airway obstruction is turned off. | ¦ | 690 | ¦ | 300 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦
¦ | Turn on airway obstruction; severity = 1.0 | ¦ | Represents the complete obstruction (severity = 1.0) of airway due to foreign body. | ¦ | 990 | ¦ | 180 | ¦<span class="success"> | Increase > Severity 0.6 @cite Morgan2006Clinical | </span>¦<span class="success"> | Increase @cite bergeronSME | </span>¦<span class="danger"> | Increase @cite bergeronSME (case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Increase @cite bergeronSME (case study of awake patient @cite izakson2013complete) | </span>¦<span class="success"> | Decrease < Severity 0.6 @cite Morgan2006Clinical | </span>¦
¦ | Remove airway obstruction | ¦ | Airway obstruction is turned off. | ¦ | 1170 | ¦ | 500 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦

## Bronchoconstriction
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Trigger Bronchoconstriction; severity 0.3 | ¦ | Trigger bronchoconstriction of severity 0.3 by Increasing effective breathing resistances on bronchi. Not severe enough to trigger CV responses. | ¦ | 30 | ¦ | 180 | ¦<span class="warning"> | Decrease @cite Morgan2006Clinical | </span>¦<span class="warning"> | Increase tachycardia @cite faust2002anesthesiology, @cite Morgan2006Clinical | </span>¦<span class="warning"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="warning"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="success"> | Decrease @cite Morgan2006Clinical | </span>¦
¦ | Remove bronchoconstriction | ¦ | Bronchoconstriction is turned off. | ¦ | 210 | ¦ | 300 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦
¦ | Trigger Bronchoconstriction; severity = 0.6 | ¦ | Trigger bronchoconstriction of severity 0.6 by Increasing effective breathing resistances on bronchi. | ¦ | 510 | ¦ | 180 | ¦<span class="success"> | Increase > Severity 0.3 @cite Morgan2006Clinical | </span>¦<span class="success"> | Increase tachycardia @cite faust2002anesthesiology, @cite Morgan2006Clinical | </span>¦<span class="success"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="danger"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="success"> | Decrease < Severity 0.3 @cite Morgan2006Clinical | </span>¦
¦ | Remove bronchoconstriction | ¦ | Bronchoconstriction is turned off. | ¦ | 690 | ¦ | 300 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦
¦ | Trigger Bronchoconstriction; severity = 1.0 | ¦ | Trigger bronchoconstriction of severity 1.0 by Increasing effective breathing resistances on bronchi. | ¦ | 990 | ¦ | 180 | ¦<span class="success"> | Increase > Severity 0.6 @cite Morgan2006Clinical | </span>¦<span class="success"> | Increase tachycardia @cite faust2002anesthesiology, @cite Morgan2006Clinical | </span>¦<span class="success"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="danger"> | Decrease hypotension @cite faust2002anesthesiology | </span>¦<span class="success"> | Decrease < Severity 0.6 @cite Morgan2006Clinical | </span>¦
¦ | Remove bronchoconstriction | ¦ | Bronchoconstriction is turned off. | ¦ | 1170 | ¦ | 500 | ¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦<span class="success"> | returns to resting physiology | </span>¦

## MainstemIntubation
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Oxygen Saturation | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Inject succinylcholine; concentration = 4820 ug/ml; Dose = 20 ml | ¦ | Neuromuscular blocker agent is administered to represent events that facilitate endotracheal intubation. | ¦ | 30 | ¦ | 60 | ¦<span class="success"> | Goes to Zero @cite Morgan2006Clinical @cite dukeSME | </span>¦<span class="warning"> | Mild Increase @cite Morgan2006Clinical 5-10% Decrease @cite dukeSME | </span>¦<span class="warning"> | Mild Increase @cite Morgan2006Clinical 5-10% Decrease @cite dukeSME | </span>¦<span class="warning"> | Mild Increase @cite Morgan2006Clinical 5-10% Decrease @cite dukeSME | </span>¦<span class="success"> | Begins to drop according to preoxygenated O2 curve | </span>¦
¦ | Turn on Anesthesia Machine and insert Endotracheal (ET) tube in the right mainstem | ¦ | Anesthesia machine is turned on to supply oxygen to the subject. ET tube is inserted, but it is set in the right mainstem. | ¦ | 60 | ¦ | 180 | ¦<span class="success"> | 16 (set by machine) | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | Continues to drop; 93-95% @cite bergeronSME | </span>¦
¦ | Reset ET tube in the trachea. | ¦ | ET tube is adjusted to be set in the trachea (good depth). | ¦ | 180 | ¦ | 480 | ¦<span class="success"> | 16 (set by machine) | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦
¦ | Turn on Anesthesia Machine and insert Endotracheal (ET) tube in the left mainstem | ¦ | ET tube is adjusted, but it is set in the left mainstem. | ¦ | 480 | ¦ | 600 | ¦<span class="success"> | 16 (set by machine) | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | No Change @cite dukeSME; Increase above normal due to cardiovascular stress @cite bergeronSME | </span>¦<span class="success"> | Continues to drop; 93-95% @cite bergeronSME | </span>¦
¦ | Reset ET tube in the trachea. | ¦ | ET tube is adjusted to be set in the trachea (good depth). | ¦ | 600 | ¦ | 900 | ¦<span class="success"> | 16 (set by machine) | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦

## EsophagealIntubation
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sample Scenario Time (s) | ¦ | Respiration Rate (breaths/min) | ¦ | Heart Rate (beats/min) | ¦ | Systolic Pressure (mmHg) | ¦ | Diastolic Pressure (mmHg) | ¦ | Oxygen Saturation (mmHg) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Administer Succinylcholine - 96.4mg | ¦ | Drug Onset < 1 minute | ¦ | 30 | ¦ | 30-50 | ¦<span class="success"> | Goes to Zero @cite Morgan2006Clinical @cite dukeSME | </span>¦<span class="danger"> | Mild Increase @cite Morgan2006Clinical; 5-10% Decrease @cite dukeSME | </span>¦<span class="danger"> | Mild Increase @cite Morgan2006Clinical; 5-10% Decrease @cite dukeSME | </span>¦<span class="danger"> | Mild Increase @cite Morgan2006Clinical; 5-10% Decrease @cite dukeSME | </span>¦<span class="success"> | Begins to drop according to preoxygenated O2 curve | </span>¦
¦ | Esophageal Intubation | ¦ | This represents failed endotracheal intubation that led to esophageal intubation | ¦ | 60 | ¦ | 120 | ¦<span class="success"> | Zero | </span>¦<span class="success"> | If undetected and O2 level drops significantly, Mild Increase @cite dukeSME | </span>¦<span class="success"> | If undetected and O2 level drops significantly, Mild Increase @cite dukeSME | </span>¦<span class="success"> | If undetected and O2 level drops significantly, Mild Increase @cite dukeSME | </span>¦<span class="success"> | continue to decrease along O2 curve | </span>¦
¦ | Successful Intubation | ¦ | Represents successful endotracheal intubation | ¦ | 180 | ¦ | 500 | ¦<span class="success"> | 16 (set by machine) | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦<span class="success"> | Towards normal @cite bergeronSME | </span>¦

## Apnea
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Tidal Volume (mL) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Apnea: Severity = 0.3 | ¦ | Mild | ¦ | 30 | ¦ | 210 | ¦<span class="success"> | Decrease to ~70% of healthy | </span>¦
¦ | Apnea: Severity = 0.0 | ¦ | Healthy | ¦ | 210 | ¦ | 510 | ¦<span class="success"> | Normal | </span>¦
¦ | Apnea: Severity = 0.6 | ¦ | Moderate | ¦ | 510 | ¦ | 690 | ¦<span class="success"> | Decrease to ~30% of healthy | </span>¦
¦ | Apnea: Severity = 0.0 | ¦ | Healthy | ¦ | 690 | ¦ | 990 | ¦<span class="success"> | Normal | </span>¦
¦ | Apnea: Severity = 1.0 | ¦ | Severe | ¦ | 990 | ¦ | 1170 | ¦<span class="success"> | Decrease to 0 | </span>¦
¦ | Apnea: Severity = 0.0 | ¦ | Healthy | ¦ | 1170 | ¦ | 1650 | ¦<span class="success"> | Normal | </span>¦